## Slide 1: Development of an AI based algorithm for detection of several detector artifacts
Vrinda Gupta
Internal Supervisor : Dr.-Ing. V. Christlein, Florian Kordon, M.Sc.
External Supervisor: Dr. Thorsten Ergler, Dr. Bastian Schmidt
In collaboration with Siemens-Healthineers

## Slide 2
Motivation
Problem Statement
Dataset and Dataset preprocessing
Experiments
  Image based
  Time based
  One-class-classifier
Comparative result analysis
Discussion and future outlook
Content

## Slide 3
Quality assurance
Manual process
Motivation
Fig: From the 3.74 million imaging tests that were taken in England in June of 2023, around 50 percent were X-ray

## Slide 4
Problem Statement
Source: https://en.wikipedia.org/wiki/Projectional_radiography
Task
Classify clinical X-ray images for artifacts
Artifact Type: Technical artifacts from a specific flat panel detector (not anatomical artifact)
Method: AI
Technical use case:
Storage of X-ray images with artifacts
Perform assessment on detectors health
Ensure high-quality diagnostic images for clinical use

## Slide 5
2350*2866 16-bit clinical raw X-ray images
Image-level labels with 7 defect and 1 good class
Detector artifacts!!
Multi-label
Dataset
Fig: MAX wi-D Detector (Source: Siemens)

## Slide 6
Artifacts
Fig: Group of Line Artifact
Fig: Partly Brighter/Darker
Fig: Sporadic Artifact
Fig: Defect Line

## Slide 7
Artifacts
Fig: Group of defect lines
Fig: Stripes

## Slide 8
Dataset
Artifact class | count
FX00_Sporadic_line_artifacts | 67
FX01_Line_artifact | 16
FX02_Group_of_line_artifacts | 6
FX03_partly_brighter | 11
FX04_Group_of_defect_lines | 9
FX05_defect_line | 74
FX09_stripes | 13
good_images | 521
Data Scarcity
Data Imbalance

## Slide 9
Dataset preprocessing
Data resampling does
not work!
2350 * 2866
1024 * 1024
Image size of 2350 * 2866 is large
Significant computational demands
Limitations on the batch size 1

## Slide 10
Dataset preprocessing
B
A
A
C
B
C
A
C
B
A- Sporadic line artifact
B- Good image
C- Partly bright/dark
2350*2866
783*955

## Slide 11
Dataset preprocessing
Original
Vertical flip
Horizontal flip
Invert
Invert Vertical
flip
Invert Horizontal
flip
Flip-both
Invert Flip-both

## Slide 12
Train - validation 80-20%
5 fold
Dataset preprocessing

## Slide 13
Motivation
Problem Statement
Dataset and Dataset preprocessing
Experiments
  Image based
  Time based
  One-class-classifier
Comparative result analysis
Discussion and future outlook
Content

## Slide 14
Image-based Approach
FX02
FX00
“7“
class
ResNet-18 (pretrained with ImageNet)
Input: Raw X-ray image patches (783,955)
Good

## Slide 15
Explainability
Fig: Artifacts and its explainabilty
Grad-CAM

## Slide 16
Inference Phase

## Slide 17
Insight: All the defects are straight line
Convert image into time series
Directional
Idea of line profile

## Slide 18
Line profiles
Sporadic
- Horizontal
Group of defect
- Horizontal/Vertical
Group of line
- Horizontal
Partly Bright/Dark
- Vertical
Stripes
- Horizontal
Defect line
- Horizontal/Vertical

## Slide 19
Rationale
Less memory requirements
Resistant to noise
No need of patching
Time series- based classification

## Slide 20
In-house developed 1DCNN model
Time series- based classification
“7“ Class

## Slide 21
Saliency Maps
Example: Defect line
Explainability
A
B
C
D
E

## Slide 22
Binary Classification
One Class Classification
Data | Count
Train | 
Good | 4684
Test | 
Good | 521
Defective | 1129

## Slide 23
One class classification
[A] Good
[B] Good
[C] FX05_Defect Line
[D] FX09_Stripes

## Slide 24
Threshold = 44.7
One class classification

## Slide 25
Often cropped; varied shape
Check model trained for intensity linear raw X-ray, works for post-processed images
Idea: Using confidence of model
Post-processed images

## Slide 26
Image based
Post-processed analysis

## Slide 27
Line Profile
Post-processed analysis

## Slide 28
Comparative Study Analysis
 | CNN-based | One class classifer | Time series-based
Works on | Image-patch | Image-patch | Line profiles
Preprocessing | Making patches of image, label each patch manually | Making patches of image, label each patch manually | Extract row and column profile and concatenate
Does classification | multiclass | binary | multiclass
Training needs | GPU | GPU | GPU or CPU
Epochs trained for | 30 | 1640 | 30
Training time(GPU) | 147 min 23 secs | 78 hours | 3min 2secs
Training time(CPU) | - | - | 14min 36sec
Works for post-processed? | Yes | - | No

## Slide 29
Comparative Study Analysis
 | Precision | Recall | F1 | Accuracy
CNN | 0.986 | 0.986 | 0.986 | 98%
1DCNN | 0.963 | 0.961 | 0.961 | 96.6%
As multi-class classification
Binary classification
 | Precision | Recall | F1 | Accuracy
CNN | 0.977 | 0.980 | 0.978 | 98.2%
OC-VAE | 0.844 | 0.847 | 0.844 | 84.4%
1DCNN | 0.968 | 0.968 | 0.968 | 96.8%

## Slide 30
Motivation
Problem Statement
Dataset and Dataset preprocessing
Experiments
  Image based
  Time based
  One-class-classifier
Comparative result analysis
Discussion and future outlook
Content

## Slide 31
Image based approach
Time Series
One class classification
Discussion
Achieved highest accuracy - 98% Works for post-processed images | Manual patching
Less computational requirement Less storage space | Works on assumption that defects are straight lines Doesn‘t seem to work directly with post-processed images
Could classify unseen artifact | High computational requirement Longer training time

## Slide 32
Cho et. al’s – ‘How much data is needed to train a medical image deep learning system to achieve necessary high accuracy?’ suggest a minimum of 4092 samples per class for achieving a 99.5% accuracy in medical imaging models
‘Context-encoding Variational Autoencoder for Unsupervised Anomaly Detection’ use 140k images of 64 × 64 size for training autoencoder to find anomaly in CT slices for achieving AUC-ROC of 0.76
Discussion

## Slide 33
Data collection and test
Data resizing exploration
Adapting to new artifact classes
Cross detector generalization
Artifact segmentation
Single label to multi-label
Future work

## Slide 34
Thank You